## Slide 1
Cold plasma endotherapy in biliary tree:
From catheter design to
Large-animal studies
Thierry DUFOUR,
Manon SOULIER,
Korentin GERAUD,
Marine CAMUS,
Laura FOUASSIER

## Slide 2
Medical context
What is cholangiocarcinoma?
Banales J.M. et al. Nat Rev Gastroenterol Hepatol 17, 557–588 (2020)
Cholangiocarcinoma = bile duct cancer
Rare: 3-60 persons for 1 billion population/y
Highly lethal: 5-year survival rate of 7-20 %
Asymptomatic in early stage
D. Moris et al., ACS Journals (2023)
Current therapies at hospital
Surgical resection Chemotherapy
20-30 %
of patients
Non curative
Surgery adjuvent
Cold Plasma: disruptive approach
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
01/22
Intra vs extra hepatic?

## Slide 3
In vitro experiments
Human cell lines of CCA (HuCCT1)
(24h post-treatment)
0 0.5 1 2 3 5
100
50
0
Plasma exposure time (min)
Cell viability (%CTRL)
In vivo experiments
Murine cell lines of CCA (SB-1)
(Immunocompetent mice)
Post-treatment time (days)
Tumor volume (mm3)
0 2 4 6 8 10
350
300
250
200
150
100
50
0
***
Plasma group
Medical context
Is cold plasma « therapeutically efficient »?
Control group
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
02/22

## Slide 4
In vivo experiments
Murine cell lines of CCA (SB-1)
(Immunocompetent mice)
Post-treatment time (days)
Tumor volume (mm3)
0 2 4 6 8 10
350
300
250
200
150
100
50
0
***
Plasma group
Medical context
Is cold plasma « therapeutically efficient »?
Control group
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
02/22
The % of tumor-infiltrating NK cells is significantly increased in the CAP-vaccine group compared with the PBS control group. Vaccination with CAP-treated tumor cells promotes the recruitment/infiltration of NK cells into the TME.
The percentage of activated NK cells is significantly higher in the CAP-vaccine group than in the PBS group. This indicates that vaccination not only increases NK cell infiltration, but also enhances their activation status.
These activated NK cells may therefore contribute more effectively to the antitumor immune response induced by CAP vaccination.
**
60
40
20
0
Tumor-infiltrating NK cells (%)
PBS CAP-Vaccine
**
1.0
0.5
0
Activated NK cells (%)
PBS CAP-Vaccine
**
60
40
20
0
Tumor-infiltrating NK cells (%)
PBS CAP-Vaccine
**
1.0
0.5
0
Activated NK cells (%)
PBS CAP-Vaccine
Vaccination
Direct effects
Indirect effects
(Vaccination with CAP-treated tumor cells)
A. Pavy & L. Fouassier

## Slide 5
Duodenoscope
Stomach
Duodenum
Pancreas
Common
bile duct
Gallbladder
How to transfer the technology from « mouse to human »?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Knowledge gap
Duodenoscope
(cross-section view)
03/22

## Slide 6
Duodenoscope
Stomach
Duodenum
Pancreas
Common
bile duct
Gallbladder
How to transfer the technology from « mouse to human »?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Knowledge gap
Duodenoscope
03/22
(cross-section view)

## Slide 7
How to transfer the technology from « mouse to human »?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Knowledge gap
03/22
Plasma
catheter
Common bile duct
Choledoc
Cystic duct
Gallbladder
+
–
–
–
–
–
–
–
–
+
+
+
+
+
+
+
Bile liquid constraints
- High electrical conductivity
- Discontinuous flow
- Surfactants
Bile ducts constraints
- Confined environment
 No gas outlet
 Overpressure
- Parietal treatment  CCA
- Axial treatment  Stenosis

## Slide 8
Plasma
catheter
Common bile duct
Choledoc
Cystic duct
Gallbladder
+
–
–
–
–
–
–
–
–
+
+
+
+
+
+
+
Bile liquid constraints
- High electrical conductivity
- Discontinuous flow
- Surfactants
Knowledge gap
Which constraints? Which challenges?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Bile ducts constraints
- Confined environment
 No gas outlet
 Overpressure
- Parietal treatment  CCA
- Axial treatment  Stenosis
Clinical device constraints
- Highly flexible
- Diameter < 4.2 mm
- Length > 2 meters
- X-rays imaging
04/22

## Slide 9
Knowledge gap
Which constraints? Which challenges?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
04/22

## Slide 10
Knowledge gap
Which constraints? Which challenges?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Catheter
Duodenoscope
Operating
channel
Guide wire
04/22

## Slide 11
Murine model
In vitro model
Post-mortem digestive porcine model
Alive
porcine
model
Patient
Artificial
biliary model
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Knowledge gap
Roadmap: upgrading preclinical models before clinical trials
05/22

## Slide 12
Copper tube
Artificial biliary model
Human bile duct
Ø 6 mm
10-3
10
10-1
10-2
1
Electrical conductivity (mS/cm)
Ø 6 mm
R
C
IR camera
Imax
Insertion tube
Proximal
control
handle
Entrance of operation channel
How to engineer an equivalent electrical model of bile duct?
Bile substitute
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Artificial biliary model
06/22

## Slide 13
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Typical voltage profile (6 kV, 5 kHz, 10%)
07/22
With bile substitute
Filamentary to-arc region
No gas breakdown

## Slide 14
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Typical voltage profile
07/22
Filamentary to-arc region
With bile substitute
No gas breakdown
No gas breakdown
Filamentary to-arc region
Without bile substitute
Typical current profile
Pulse FWHM
= 150 ns

## Slide 15
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
08/22
Filamentary to-arc region
With bile substitute
No gas breakdown
CEI 60479-2
CEI 60479-1
Technical specifications
(IEC, International Electrotechnical Commission)

## Slide 16
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
08/22
Filamentary to-arc region
With bile substitute
No gas breakdown
International Electrotechnical Commission (IEC) standards
Current magnitude (A)
Pulse duration (ms)
Our
condition
c1
c2
c3
DC-1
DC-2
DC-3
b
DC-4.3
DC-4.2
DC-4.1
Slight tingling sensation possible
Involuntary muscle contractions
Strong involuntary muscle contractions
Probability of ventricular fibrillation: 0-5%.
Probability of ventricular fibrillation: 5-50%.
Probability of ventricular fibrillation: >50%.
Our experimental conditions !!
RMS body current(A)
Duration of 1 current pulse (s)

## Slide 17
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
09/22
Filamentary to-arc region
No gas breakdown
Sequence of N pulses
Parameters
Current pulse duration
RMS body current
Pulse repetition frequency
Rest time between 2 current pulses (2pulses)
2pulses  20 µs
If 2pulses < Cardiac period (0.86s)
Cumulative effects: at each new pulse, the fibrillation roof decreases  Risks are enhanced.
Ventricular
roof
fibrillation
2pulses

## Slide 18
Artificial biliary model
Defining a safe operational window / Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
09/22
Filamentary to-arc region
No gas breakdown
Sequence of N pulses
Parameters
Current pulse duration
RMS body current
Pulse repetition frequency
Rest time between 2 current pulses (2pulses)
2pulses  20 µs
If 2pulses < Cardiac period (0.86s)
Cumulative effects: at each new pulse, the fibrillation roof decreases  Risks are enhanced.
Our plasma physics goes beyond IEC standards
Current pulse durations are too short
Alternance of positive & negative pulses

## Slide 19
Artificial biliary model
Defining a safe operational window / Temperature
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
10/22
Filamentary to-arc region
No gas breakdown
Plasma
ON
Plasma
OFF
Plasma
OFF
(Without bile substitute)
INFRARED IMAGING
Data
- Emissivity (ε): 0–1.
- Black/matte surfaces (matte paints, black electrical tape) ≈ 0.95.
- Polished metals ≈ 0.02–0.2.
If the correct emissivity is not enterred, the error can be tens of degrees (polished metals will “look cold” if ε = 0.95 is assumed, whereas they’re typically ε ≈ 0.02–0.2).
Relations
Trends
- Oxidation/roughening → ε increases.
- Material shiner  the more it reflects the environment  Raw reading is more biased.
- If you enter an ε that’s too high, the camera underestimates Treal; if it’s too low, it overestimates Treal.
Tapp : Apparent temperature alias radiance received by the IR camera
Tref: temperature (or radiance) from the surroundings reflected by the target.
1/ place a small crumpled-then-flattened aluminum foil patch at the target location, same view/angle. The crumpled foil has very low ε and behaves as a diffuse mirror in LWIR  its Tapp corresponds to the environment’s temperate rather than its own emission.
2/ Set camera emissivity = 1.00 (or ~0.98) and distance small, then read the foil’s apparent temperature.
3/ Enter that value as Reflected Apparent Temperature (RTC/RAT) in the camera; it internally converts to Tref for its band.
How to measure correctly
Enter the right ε for the camera’s spectral band (typically 8–14 µm).
Correct for reflected radiation (reflected apparent temperature, RAT/RTC): measure it using a crumpled aluminum-foil target viewed from the same position, then input that value.
Prepare the surface if ε is low:
Apply a matte black tape/paint spot (ε ≈ 0.95) and measure on it; or
Slightly roughen/oxidize the surface to raise ε.
Viewing angle: stay near normal incidence (< 45°); emissivity drops with angle on smooth surfaces.
Distance/atmosphere: if available, correct for distance, humidity, and air temperature (important over long paths or hot hoods).
Transparency: some plastics/IR windows aren’t opaque in the band; beware of windows (ordinary glass is not transparent in 8–14 µm).
Contact check: to calibrate a material’s ε, compare with a thermocouple/contact probe and adjust ε until they match.
(And to your other point: doing measurements “in the dark” in visible light does not fix this—surfaces still reflect surrounding thermal IR from walls, equipment, and people.)

## Slide 20
Artificial biliary model
Defining a safe operational window / Temperature & Electricity
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
10/22
Plasma
ON
Plasma
OFF
Plasma
OFF
(Without bile substitute)
Filamentary to-arc region
With bile substitute
T > 40 °C
No gas breakdown
Filamentary to-arc region
Without bile substitute
T > 40 °C
Safe Operating Window: 5-7 kV & 5-10 kHz
No gas breakdown

## Slide 21
Submerged catheter
Catheter with 4mm liquid gap
Worst scenario
No interaction with gas phase
H2O2 significant production
Treatment time = 10 min
No plasma
5 kHz – 6 kV
10 kHz – 6 kV
Treatment time = 10 min
No plasma
5 kHz – 6 kV
10 kHz – 6 kV
More realistic scenario
Interaction with gas phase
Increase in RONS production
Artificial biliary model
Antitumoral potential: RONS concentration in liquid
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context

## Slide 22
PELP
(Plasma EndoLuminal Procedure)
Digestive porcine model
A mandatory milestone to refine plasma parameters
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Digestive porcine model
Endoscopy column
Plasma catheter
Duodenoscope
Post-mortem digestive model
11/22

## Slide 23
Post-mortem digestive model
Defining the safe operating window
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Electrical measurements (6 kV, 5 kHz, 10%)
Thermal measurements (6 kV, 5 kHz, 10%)
12/22
B
C
B
A
C
Maximum current (mA)
T(t) – Tt=0min (°C)
Time (min)

## Slide 24
N = 7 pigs
N = 4 pigs
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Euthanasia
Autopsy
Histological
analysis
Context
Hôpital Européen Georges Pompidou (HEGP)
4 experimental campaigns
18 healthy domestic pigs weighing 25-30kg
  7 pigs: optimization
  11 pigs: replication
Euthanasia
Can we endoscopically deliver plasma into porcine bile duct?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
Short-term side effects?
Long-term side effects?
13/22
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)

## Slide 25
Alive porcine model
Can we endoscopically deliver plasma into porcine bile duct?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
14/22
Plasma EndoLuminal Procedure (PELP)
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)
Plasma-assisted endoscopic retrograde cholangiopancreatography
Porcine model
X-ray imaging system
Endoscopic column
George Pompidou European Hospital (Paris). Marine Camus, MD in digestive endoscopy, St-Antoine Hospital (Paris)

## Slide 26
Plasma EndoLuminal Procedure (PELP)
Plasma-assisted endoscopic retrograde cholangiopancreatography
Porcine model
X-ray imaging system
Endoscopic column
George Pompidou European Hospital (Paris). Marine Camus, MD in digestive endoscopy, St-Antoine Hospital (Paris)
Duodenal
papilla
Duodenoscope camera
Alive porcine model
Can we endoscopically deliver plasma into porcine bile duct?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
15/22

## Slide 27
Challenge #1: Localize the duodenal papilla and enlarge it with ad hoc endoscopic tools
Duodenal
papilla
Duodenal
papilla
Duodenoscope camera
Alive porcine model
The 4 challenges we had to face with …
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
15/22

## Slide 28
Challenge #1: Localize the duodenal papilla and enlarge it with ad hoc endoscopic tools
Challenge #3: Find a balance between gas flow rate, gas injection time and pressure
Challenge #2: Insert the plasma catheter in the papilla
Plasma catheter
Duodenal
papilla
Duodenal
papilla
Duodenoscope camera
Alive porcine model
The 4 challenges we had to face with …
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
15/22

## Slide 29
Without contrast agent
With contrast agent
X-rays
Guidewire
Plasma
catheter
Endoscope
Common
bile duct
Endoscope
Guidewire
& catheter
Alive porcine model
The 4 challenges we had to face with …
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
No more camera in the biliary tree  Challenge #4: How to navigate the plasma catheter?
15/22

## Slide 30
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
16/22
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Euthanasia
Autopsy
Histological
analysis
Euthanasia
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)

## Slide 31
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
16/22
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
7 kV, 5 kHz, 10%, He flow rate = 0.7 slm, N=11
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)

## Slide 32
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
16/22
Max
Min
Max
Min
Before During After
Before During After
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)

## Slide 33
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
16/22
Cholangiograms
After PELP
Before PELP
No perforation of the biliary tract
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma therapy: 5 min
Plasma EndoLuminal Procedure (PELP)

## Slide 34
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
17/22
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
1 cm
1 cm
Duodenal papilla
Choledoc
Gallbladder
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma EndoLuminal Procedure (PELP)
Plasma therapy: 5 min

## Slide 35
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
18/22
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma EndoLuminal Procedure (PELP)
Plasma therapy: 5 min
200 µm
PELP Biliary tissue 30 min later
100 µm
SHAM biliary tissue, 30 min later
 | Sham biliary tissue | Plasma-treated biliary tissue
Luminal epithelium |  Mostly continuous and preserved biliary epithelial lining.  No epithelial denudation or ulceration. |  Largely preserved, but the luminal surface appears more irregular, folded, and reactive.  No obvious full-thickness epithelial loss.
Subepi-thelial stroma | Pink fibroconnective layer, with scattered nuclei and small vessels. |  Possible mild inflammatory /vascular changes.  Some congestion can be observed, suggesting a limited tissue response after plasma.
Periductal connective tissue | Preserved periductal architecture with gland-like structures, vessels, and connective tissue. |  Overall architecture remains recognizable.  No obvious transmural destruction or necrosis
Adipose tissue | Deep layer composed of clear vacuolated adipocytes surrounding the biliary wall. | Structurally preserved. No obvious plasma-induced damage.
Largely preserved even if more irregular, folded & reactive.
Mild vascular congestion / inflammatory reaction.
Architecture preserved. No obvious transmural destruction or necrosis
Structurally preserved. No obvious plasma-induced damage.
Luminal epithelium
Subepithelial stroma
Periductal connective tissue
Adipose tissue

## Slide 36
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma EndoLuminal Procedure (PELP)
Plasma therapy: 5 min
40 days of monitoring
(Veterinary supervision: staff of Georges Pompidou European Hospital)
Social behavior
Feeding
Health
No pain signs
Well-being
Liveliness
All parameters remained consistent with a positive welfare profile.
How about potential side effects?
Alive porcine model
19/22
Liver disorder ?
Cell damage?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context

## Slide 37
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
20/22
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma EndoLuminal Procedure (PELP)
Plasma therapy: 5 min
Day 1
Day 40
No macroscopic damage

## Slide 38
Animal acclimatization period
(10 days)
40days
of monitoring
including blood
samplings
Autopsy
Histological
analysis
Autopsy
Histological
analysis
Euthanasia
Euthanasia
Anesthesia: 1-2h
Gas delivery: 10-30 min
Plasma EndoLuminal Procedure (PELP)
Plasma therapy: 5 min
How about potential side effects?
Knowledge
gap
model
model
Post-mortem digestive
Porcine
model
Artificial biliary
Medical
context
Alive porcine model
21/22
200 µm
PELP: biliary tissue 40 days later
100 µm
SHAM: biliary tissue 40 days later
Compared with SHAM tissue  Increased of thickening of the subepithelial/periductal stroma  Mild chronic reparative or fibrotic response
 No obvious ulceration, transmural necrosis, hemorrhage or deep structural destruction

## Slide 39
STRENGTHS
WEAKNESSES
Murine models: antitumor effects demonstrated
Porcine models: electrical & thermal risks controlled under bile flow constraints  No side effects
Plasma catheter fully matches with conventional endoscopic material
Combo : plasma (local) + immunotherapy (systemic)
2 patents, maturation starting in 3 months (certification & pre-industrial catheter)
Clinical trials expected in 2 years
Spinoffs: pulmonary cancer (NSCLC)
OPPORTUNITIES
THREATS
Trade-off: therapeutic efficacy vs Anesthesia
Shortage & volatility of helium gas
From bench to bedside: will preclinical antitumor efficacy be the same on patients?
Alive porcine models
  Chemical environment unknown  Which ROS are plasma-generated?
  In situ & real time measurements of plasma parameters remains challenging
EMR
Conclusion & Outlooks
22/22

## Slide 40
Thank you for your attention
17/22